## Sheet1
TEST SYSTEM: Curian Campy
Steps | Failure Mode | Causes | Means of Failure Detection
|  |  | Internal Controls | External Controls | Engineering Controls | Operator Training | Laboratory's Mitigating Actions
List the stage or aspect of the test system’s process under investigation. | List all manners in which failure could occur in this step. | List all causes of the failure mode that have the potential to produce incorrect test results. | Are there internal biological or procedural controls to detect failure? | Can external controls increase the probability to detect failure? | Are there manufacturer checks to reduce the probability of failure? | Can operators reduce or detect failures through training? | What other processes can the laboratory implement to reduce failure?
Specimen collection | Specimen degradation or specimen contamination | Improper Specimen type | No | No | No | Yes | Techs are trained that only unpreserved stool or stool in Cary Blair media is the approved specimen type. Techs will reject all other specimen types.
|  | Improper Collection | No | No | No | Yes | Techs will reject all non-stool specimens, including colonoscopy and colostomy specimens.
|  | Improper storage and transport | No | No | No | Yes | Laboratory will provide collection, storage and transportation instructions to both patients and clinicians.
|  | Interfering substances | Yes | No | Yes | No | No interfering substances were identified in manufacturer's studies. Internal control must develop and be recognized by the instrument for a valid result to occur.
Specimens and Reagents | Sample Processing | Insufficient mixing of stool specimen | No | No | No | Yes | 1. Training and competency assessment must be completed and signed off prior to performing any testing where patient results are reported.
2. Competency is completed annually
3. Each user participates in proficiency testing annually
|  | Incorrect amount of stool added to Aioprep device | Yes. Device has internal flow control. | No | Yes | Yes | Competency of all users is assessed annually. All users will participate in proficiency testing annually.
|  | Incorrect Sample Collection brush used | No | No | Yes, brushes are color coded to match Aioprep device | Yes | 1. Training and competency assessment must be completed for each user prior to performing patient testing
2. Competency of all users is assessed annually.
3. Eash user participates in proficiency testing annually.
|  | Specimen or kit not warmed to room temperature prior to testing | Yes, internal control must develop and be detected by analyzer for valid result. | No | Yes, internal control must develop and be detected by analyzer for valid result. | Yes | 1. Training and competency assessment must be completed for each user prior to performing patient testing
2. Competency of all users is assessed annually.
3. Each users participates in proficiency testing annually.
|  | Incorrect amount of prepared specimen added to test card | Yes, internal control must develop and be detected by analyzer for valid result. | No | Yes, internal control must develop and be detected by analyzer for valid result. | Yes | Competency of all users is assessed every six months. All users will participate in proficiency testing annually.
|  | Incorrect test card used | No | No | Yes, instrument recognizes barcode etched onto test card and reports result according to test card used. | Yes | Competency of all users is assessed every six months. All users will participate in proficiency testing annually.
|  | Incorrect analyze option chosen on instrument | Yes, if Analyze Now option is selected and card has not been incubated, internal control would not be developed and test would be considered invalid. | No | Yes, if Analyze Now option is selected and card has not been incubated, internal control would not be developed and test would be considered invalid. | Yes | Competency of all users is assessed every six months. All users will participate in proficiency testing annually.
Operator | Operator Capacity | Untrained operator | Yes, internal control monitors for adequate sample preparation | Yes, untrained operators may not obtain correct external control results | Yes, Admin user may configure instrument so that only qualified users may log in and use equipment | No | Admin user may configure instrument so that only qualified users may log in and use equipment. If needed, daily QC and instrument check tests may be required for testing.
| Laboratory Staffing | Short staffing | No | No | No | No | In the event of short staffing, TAT may be impacted. Only trained operators can run the test and specimens may be run in larger batches.
|  | Underqualified staff | No | No | No | No | Hiring manager and human resources will verify all operators are qualified per local, federal and state guidelines for moderately complex testing.
Reagents | Test Kit Degradation | Shipping | Yes, internal controls to detect reagent integrity | Yes, external controls monitor kit and reagent degradation. | Yes, valid internal control required for reportable results. | Yes | Product is shipped in labelled boxes that detail required storage conditions. Operators will notify Laboratory Manager if there are any defects are noticed upon receipt of kit or it kit is failing produce valid internal or external control results.
|  | Storage | Yes, internal controls to detect reagent integrity. | Yes, external controls monitor kit and reagent degradation. | Yes, valid internal control required for reportable results. | Yes | Laboratory SOP will document proper storage of kit.
|  | Used past expiration | Yes, internal controls to detect reagent integrity | Yes, external controls monitor kit and reagent degradation. | Yes, valid internal control required for reportable results. | Yes, operators are trained to check test kit expiration dating prior to initiating testing | Expiration date checked and documented with each run
| Curian analyzer failures | Optics | Yes, valid internal control must be recoginized by optics with each test card analyzed to produce valid tet results | Yes, external control material can be used to monitor function of analyzer including optics | Yes, power on self test and instrument check test are both used to monitor optic function | No | Instrument check test will be run every #### days at this facility
|  | Test card drawer not fully closed | Yes, internal control bands must be properly located inside of instrument to be considered valid | No | Yes, instrument has sensor to detect if drawer is closed prior to initiating test. | Yes | Ensure staff is trained to proper use of instrument.
|  | Software failures | No | No | Yes, power on self test validates function of software prior to use. | No | Ensure most up to date software is installed on unit.
Laboratory Environment | Atmospheric Environment | Dust | No | No | Yes, instrument check test may fail due excess dust on optic system | No | Daily cleaning of benches is required for both dust and microbial contamination.
|  | Temperature | Yes, out of range room temps may affect internal control line formation | Yes, out of range temps may cause external control failures. | No | Yes | Temperature of the lab is monitored and documented daily.
Laboratory Environment |  | Humidity | No | No | No | Yes | Humidity is monitored by the lab. Equipment will not be used if humidity is outside of manufacturer specifications.
| Utility Environment | Electrical | No | No | Yes, power on self test verifies proper electrical requirements are met . Failed test locks operators from using instrument. Instrument also requires installation of ferrites to prevent electrical interference. | No | Eletrical requirements for all lab is maintained by facilites department.
|  | Water Quality | N/A | N/A | N/A | N/A | N/A
|  | Air Pressure | N/A | N/A | N/A | N/A | N/A