
[image: ]

Training Checklist
 (Please see Package Insert for complete instructions)

	
	Name _____________________________________ Facility _________________________________


	
	Kit and Specimen collection, handling, and storage


	|_|
	[bookmark: _Hlk97193555]Trainee understands kits should be stored at 2-8 C when not in use and allowed to warm up to room temperature for 30-60 minutes prior to use.  

	|_|
	Trainee understands specimens should be brought to room temperature prior to testing. 


	[bookmark: Check1]|_|
	Trainee understands stool specimen should be received unpreserved or preserved in Cary-Blair (C&S) type media. 

	|_|
	Trainee understands unpreserved stool specimens can be stored 2-8 C for 7 days or ≤ -20 C for 8 weeks.

	|_|
	Trainee understands that preserved stool specimens can be stored at 2-8 C 7 days, frozen ≤-20 C 8 weeks and 19-27 C (RT) for 96 hours.

	|_|
	Trainee understands that all specimens to be tested must be thoroughly mixed by vortexing (liquid-semi-solid) or mixed with a wooden stick for formed/solid specimens.


	|_|
	For preserved stool specimens (Cary-Blair or C&S), trainee understands how to process 300 ul of preserved stool using a provided transfer pipette into the Aioprep device.


	|_|
	For unpreserved stool specimens, trainee understands how to process 100 ul unpreserved stool into the Aioprep device.

	|_|
	Trainee understands that non-pipettable formed/solid specimen whether in Cary-Blair (C&S) media or unpreserved should be transferred into the Aioprep device using a collection brush.
Curian Campy brushes are only available by request through Technical Support. 

	|_|
	Trainee understands that specimens inoculated Aioprep should be vortexed for approximately 5 seconds and never shaken or inverted.

	|_|
	Trainee understands that specimens in the Aioprep device may be held at RT (19-27 C) for 8 hours or at 2-8 C for up to 24 hours.

	|_|
	Trainee understands that to inoculate the test device, they need to unscrew the white bottom cap, hold the
Aioprep vertically and in the center of the tube and squeeze 3 drops of sample into the sample port of the test card. Then discard the white cap and Aioprep tube.

	|_|
	Trainee understands to immediately navigate to the Test menu on the analyzer and select ANALYZE NOW or INCUBATE and ANALYZE and follow any on screen directions.


	
	Quality Control & Interpretation of Results


	|_|
	Trainee understands Manufacturer recommends external positive and negative controls should be run per kit lot / shipment. 

	[bookmark: Check7]|_|
	Trainee understands to run the PC add 3 drops of Positive Control directly to test card
Trainee understands to use the Aioprep diluent as the Negative Control adding 3 drops from the device onto the test card

	|_|
	Trainee understands to load both the Positive and Negative control cards onto the Curian device and choose either the ANALYZE NOW or INCUBATE and ANALYZE program following screen instructions.

	[bookmark: Check8]|_|
	Trainee understands that if using the ANALYZE NOW procedure, the test card needs to be read on the Curian analyzer within 2 minutes following the 20-minute incubation on the bench.

	[bookmark: Check9]|_|
	Trainee understands that the kit should not be used if control tests do not produce the correct results.

	
	Precautions and Limitations


	[bookmark: Check14]|_|
	Trainee understands that transferring too little specimen or failure to mix and completely suspend the specimen in the sample diluent of the Aioprep may result in a false negative result.

	[bookmark: Check15]|_|
	Trainee understands that non-pipettable, formed/solid should only be processed with the Curian Campy Sample Collection brush 



Trainee:__________________________________________________________Date:________________

Trainer:__________________________________________________________Date:________________
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